LODE · Custom Session SOP JEOL JSM-IT800 SEM-EDS — fracture surface morphology Booking VEIN-0006 · Generated 2026-05-25 14:55

INSTRUMENT JEOL JSM-IT800 SEM-EDS; RESEARCHER Marcus Rivera; RESEARCH GROUP Steel embrittlement; MATERIAL martensitic steel. SAMPLE 5mm x 5mm; SESSION Tue, May 26 · 13:00–16:00 EDT; PREP STARTS Tue, May 26 · 12:30; FIT SCORE 92/100 (A).

CUSTOMIZED SAMPLE-SPECIFIC NOTES This procedure is optimized for Marcus Rivera's 5mm x 5mm carbon-coated martensitic steel sample to analyze fracture surface morphology under high-urgency conditions before the Thursday deadline. Special attention is paid to mitigating the edge charging issues previously observed in Marcus's prior runs on this instrument.

RECOMMENDED OPERATING PARAMETERS

| Parameter | Recommended value (with citation) |
| --- | --- |
| Accelerating Voltage | 15 kV [[1] Historical Run Log, Run #2] |
| Sample Dimensions | 5mm x 5mm [Booking Context] |
| Coating Status | Carbon coated [[1] Historical Run Log, Run #2] |
| Primary Imaging Mode | Secondary Electron (SE) Imaging [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112] |
| Analytical Method | EDS Mapping [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112] |

PRE-SESSION CHECKLIST  (CUSTOMIZED TO SAMPLE)

☐  Verify the sample is securely mounted and carbon-coated to prevent edge charging.  [[3] Historical Run Log, Run #6]

☐  Confirm the JEOL JSM-IT800 chamber is vented and ready for sample loading.  [[7] jeol-sem, jeol-sem]

☐  Check that the EDS detector is cooled and fully operational.  [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112]

☐  Set initial stage height to a safe working distance for a 5mm sample.  [[8] mines-sem-sop, mines-sem-sop]

☐  Ensure the user is logged into the Mines SEM booking system.  [[8] mines-sem-sop, mines-sem-sop]

STEP-BY-STEP PROCEDURE

1. Sample Mounting and Loading — Mount the 5mm x 5mm carbon-coated martensitic steel sample onto the stage stub, ensuring excellent electrical contact to mitigate edge charging during high-voltage analysis.  [[3] Historical Run Log, Run #6]

2. Chamber Evacuation — Insert the sample holder into the JEOL JSM-IT800 chamber and initiate the pump-down sequence to achieve high vacuum.  [[7] jeol-sem, jeol-sem]

3. Beam Initialization — Turn on the electron gun and gradually increase the accelerating voltage to 15 kV, aligning the beam for optimal resolution.  [[1] Historical Run Log, Run #2]

4. Focus and Astigmatism Correction — Locate the polished fracture surface, focus at low magnification, and correct astigmatism to ensure sharp imaging of the morphology.  [[8] mines-sem-sop, mines-sem-sop]

5. Secondary Electron Imaging — Acquire high-resolution SE images of the fracture surface morphology to evaluate embrittlement features and grain boundary decohesion.  [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112]

6. EDS Mapping Acquisition — Position the EDS detector and initiate elemental mapping at 15 kV to analyze grain boundary and precipitate chemistry.  [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112]

7. Charging Mitigation Monitoring — Monitor the edges of the martensitic steel sample closely; if charging is observed, adjust the scan speed or probe current immediately.  [[3] Historical Run Log, Run #6]

POST-RUN CHECKLIST

☐  Turn off the electron beam and retract the EDS detector safely.

☐  Vent the JEOL JSM-IT800 chamber and carefully retrieve the 5mm x 5mm steel sample.

☐  Pump down the chamber to leave the instrument in a high-vacuum state.

☐  Log off the instrument control software and clean the workspace.

SAFETY NOTES

•  Always wear gloves when handling the sample holder and chamber interior to prevent contamination.  [[8] mines-sem-sop, mines-sem-sop]

•  Ensure the EDS detector is fully retracted before moving the stage to prevent collision.  [[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide, 6.1 EDS vs EBSD Selection, p.112]

•  Do not exceed the maximum sample height limits to avoid damaging the pole piece.  [[7] jeol-sem, jeol-sem]

RAG REFERENCES

This SOP was generated by LODE's authoring agent grounded in the following lab knowledge sources (manuals, lab SOPs, maintenance logs).

[1] Historical Run Log — Run #2

[2] Historical Run Log — Run #8

[3] Historical Run Log — Run #6

[4] Historical Run Log — Run #7

[5] Historical Run Log — Run #5

[6] JEOL JSM-IT800 EDS/EBSD Attachment Guide — 6.1 EDS vs EBSD Selection (p.112)

[7] jeol-sem — jeol-sem

[8] mines-sem-sop — mines-sem-sop

Fit rationale: SEM with EDS is optimal for fracture surface chemistry mapping at the sample scale described. [Source: JEOL JSM-IT800 Manual, Section 3.2, p.42]