LODE · Custom Session SOP Bruker D8 Advance XRD — phase identification Booking VEIN-0020 · Generated 2026-05-25 17:36

INSTRUMENT Bruker D8 Advance XRD; RESEARCHER Test John; RESEARCH GROUP test; MATERIAL chalcopyrite. SAMPLE —; SESSION Tue, May 26 · 09:00–11:00 EDT; PREP STARTS Tue, May 26 · 08:15; FIT SCORE 78/100 (B+).

CUSTOMIZED SAMPLE-SPECIFIC NOTES This SOP is customized for Test John's phase identification on chalcopyrite. Cited RAG sources are listed at the end of this document.

RECOMMENDED OPERATING PARAMETERS

| Parameter | Recommended value (with citation) |
| --- | --- |
| Radiation | Cu Kα [Historical Run Log Run #1] |
| 2θ range | 10–80° [Historical Run Log Run #1] |
| Step size | 0.02° [Historical Run Log Run #1] |
| Count time | 2 s / step [Historical Run Log Run #1] |
| Tube current | ≤ 40 mA [bruker-d8-xrd bruker-d8-xrd] |

PRE-SESSION CHECKLIST  (CUSTOMIZED TO SAMPLE)

☐  Confirm safety training for Bruker D8 Advance XRD is current

☐  Arrive by 08:15 for warm-up and prep

☐  Review SDS for chalcopyrite and lab EH&S guidelines  [bruker-d8-xrd bruker-d8-xrd]

☐  Bring USB drive labelled with research group

STEP-BY-STEP PROCEDURE

1. Warm-up — Power on the instrument and complete its warm-up cycle (~20 min).  [bruker-d8-xrd bruker-d8-xrd]

2. Load sample — Mount chalcopyrite (as supplied). Verify surface level and stage clearance.  [Historical Run Log Run #1]

3. Configure — Apply parameters above for phase identification.  [Historical Run Log Run #1]

4. Acquire data — Run the planned scan / image series for phase identification. Monitor for saturation or charging.

5. Export — Save raw files to lab server under your group folder.

6. Shutdown — Follow instrument-specific shutdown checklist in the lab binder.

POST-RUN CHECKLIST

☐  Submit LODE post-run report with actual parameters + quality rating

☐  Log instrument hours and any anomalies in the maintenance log

☐  Backup raw data to research group share within 24 hours

☐  Notify lab manager of any DET-SAT, VAC-LOW, or charging issues

SAFETY NOTES

•  Handle chalcopyrite per Mines EH&S guidelines.  [bruker-d8-xrd bruker-d8-xrd]

•  PPE: safety glasses (and gloves when handling powders).

RAG REFERENCES

This SOP was generated by LODE's authoring agent grounded in the following lab knowledge sources (manuals, lab SOPs, maintenance logs).

[1] Historical Run Log — Run #1

[2] bruker-d8-xrd — bruker-d8-xrd

[3] Bruker D8 Advance Manual — 4.3 Phase Identification (p.67)

[4] XRD Maintenance Log — DET-SAT History

[5] Historical Run Log — Run #4

[6] Mines XRD Lab SOP — Safety and Warm-up (p.1)

[7] Bruker D8 Advance Manual — 5.1 High-Strength Steel Analysis (p.89)

Fit rationale: XRD suitable for phase identification; use Co Kα for steel to avoid fluorescence. [Source: Bruker D8 Advance Manual, Section 5.1, p.89]